## Instructions
NTC Thermistor Calibration Template — How to Use
Purpose
Fit two models to each thermistor: the 2-parameter B model (R₀, B) and the 3-parameter Steinhart-Hart model (A, B, C). Plot of ln(R) vs 1/T should be a straight line with slope B. The Steinhart-Hart fit captures the mild curvature the B-model can't.
Procedure — thermal setup
1. Fill a small insulated water bath (a vacuum flask or stirred beaker on a hotplate works). Bond the two thermistors together with your reference probe (multimeter temp probe or DS18B20) using a small blob of thermal compound wrapped in thin heat-shrink or PTFE tape. Goal: identical thermal mass touching the same water. 2. Stir continuously — unstirred water can have several-°C gradients between the probe and the thermistor only centimetres away. A small aquarium pump or magnetic stirrer is perfect. 3. Start cold (ice slurry, not just cold tap water — crushed ice + water gives a rock-solid 0.0 °C plateau for ~10 minutes). Then work upward: 10, 20, 30, 40, 50, 60, 70, 80 °C. You can drop points if you're short on time — eight good points is plenty. 4. At each set-point, wait until BOTH the reference reading and the thermistor resistance stop drifting (usually 60–120 s after you stop adding heat or ice). Record T_ref and R_meas simultaneously into the Thermistor_1 / Thermistor_2 sheets.
Procedure — electrical
• Measure R in 2-wire mode on your multimeter if R is above ~1 kΩ; use 4-wire if it's lower. • Keep your DMM's test current low (most meters apply <1 mA on the kΩ range); glass-bead NTCs have dissipation constants around 1–2 mW/°C in still water, much better when submerged and stirred. At 1 mA through a 10 kΩ bead, P = 10 mW, which gives a few mK of self-heating in stirred water — fine. In still air you'd get several °C of error. • Leave the same wiring between bench and bath throughout calibration. Contact resistance changes are your enemy.
What the columns do
T_ref (°C) and R_meas (Ω) are your inputs (blue cells). Everything else auto-computes: T in Kelvin, 1/T, ln(R), the fitted resistance from the B-model, and the residual %. The (ln R)³ helper columns feed the LINEST call that solves the Steinhart-Hart coefficients.
Reading the fit
• B value: typical glass-bead NTCs are 3000–4500 K. A value outside that range usually means either the part is unusual or your data has a bad point. • R²: you want ≥ 0.999 for a clean bead. Lower than 0.995 means you've got non-equilibrium points or a bad contact. • Residual (%): should be within ±1 % across most of the range for the B-model; the Steinhart-Hart fit typically brings this down to <0.1 % if your data is clean.
Google Sheets notes
This file opens directly in Google Sheets (File → Import → Upload). All formulas (SLOPE, INTERCEPT, RSQ, LINEST, INDEX, IFERROR, EXP, LN) work identically there. The embedded scatter chart may render as a static object; if you want to edit it, right-click → 'Open with Google Sheets' or recreate: Insert → Chart → Scatter → X = column E (1/T), Y = column F (ln R).

## Analysis Notes
Analysis Notes — v2 additions
What changed
Added a new analysis block at the bottom of each thermistor sheet (rows 45–66) and six new helper columns (L–Q). The original 20-point full-range B-model and Steinhart-Hart fits are untouched, so you can compare the two approaches side by side. The Comparison sheet has a new block of rows (15 onward) showing the narrow-range results.
Why a narrow-range fit was added
The original full-range SH fit has ~0.27 °C residual RMS because of non-equilibrium data at the cold (points 1–4) and hot (points 18–20) extremes — the bath was ramping, not settled. These outer points dragged the global fit coefficients around even though they won't be visited during operation. Fitting only the 5–40 °C subset (10 points per thermistor) gives coefficients that are specifically tuned to the quiet middle of the calibration, and residuals in the 10–30 °C operating range collapse to ~0.10 °C RMS — close to what your reference probe's 0.1 °C digit can even resolve.
Why B-model and not Steinhart-Hart
The classic curvature that justifies Steinhart-Hart (the "smile" residual pattern) only matters when you fit across a wide temperature span. Over a 35 °C fit window (5–40 °C), the B-model — a straight line of ln R versus 1/T — is effectively indistinguishable from SH. Fewer parameters also means less risk of overfitting, and the two-parameter form is trivial to invert in firmware: T = B / (ln R − intercept) − 273.15. If you later widen the operating range beyond ~40 °C, re-evaluate.
How the fit is calculated (since SLOPE/LINEST wouldn't work)
SLOPE, INTERCEPT, and LINEST refuse ranges that contain non-numeric values, and our in-range mask leaves zeros for ignored rows. So the B-fit is computed by hand via the closed-form least-squares formulas: slope = (n·Σxy − Σx·Σy) / (n·Σxx − (Σx)²) and intercept = (Σy − slope·Σx) / n, where x = 1/T_K, y = ln R, and the sums run over rows where the mask indicator L = 1. R² is computed from the same sums. This is exactly what SLOPE/INTERCEPT do internally; we just made the masking explicit.
How to tune the fit
The four blue input cells (B47–B50) control the two windows. B47–B48 define which points enter the fit (default 5 to 40 °C). B49–B50 define which points are used for the residual statistics (default 10 to 30 °C, your actual operating range). Change either pair and every dependent cell recomputes. Useful sanity checks: tightening the fit window to 10–30 °C gives the purest operating-range fit but fewer degrees of freedom (6 points, 4 DOF); widening to 0–60 °C starts to let the bad tail points back in and inflates residuals.
What to do with the narrow-range coefficients
Paste the narrow-range B (row 53) and intercept (row 54) from each thermistor sheet into the Arduino firmware's per-channel calibration. The conversion is R → T using T_C = B / (LN(R) − intercept) − 273.15. If you'd rather work with the Steinhart-Hart form (A + B·ln R + C·(ln R)³), the existing full-range SH coefficients at rows 40–42 are still usable — just expect the ~0.27 °C residual RMS rather than ~0.10 °C. For a precision-only application in 10–30 °C, the narrow B-fit is the better choice.
Helper columns (L–Q) reference
L = in-fit-range indicator (1 if B47 ≤ T_ref ≤ B48, else 0). M = in-op-range indicator. N = 1/T in kelvin, but zeroed when L=0 (so SUMPRODUCT works). O = ln R, zeroed when L=0. P = temperature predicted from measured R via the narrow-range fit (shown for every data point so you can spot-check). Q = residual in °C, but blank outside the operating range so MIN/MAX/COUNT/SUMSQ aggregate only over the relevant subset.
Matching between the two thermistors (updated)
Full-range fits: B differs by 15.7 K (0.41 %), R₀ by 631 Ω (0.64 %). Narrow-range fits: B differs by ~4 K (0.1 %), R₀ by ~340 Ω (0.3 %). The pair is more closely matched than the original fits suggested — the extra scatter in the wide-range fits was absorbing non-equilibrium error, not real bead-to-bead variation.

## Thermistor_1
NTC Thermistor Characterization — NTC-01
Thermistor ID: | NTC-01
Reference T₀ (°C): | 25 | T₀ (K): | 298.15
Measurement Data (fill the blue cells — T_ref is your reference probe, R_meas is multimeter reading)
Point # | T_ref (°C) | T_ref (K) | R_meas (Ω) | 1/T (1/K) | ln(R) | R_fit B-model (Ω) | Residual (%) | (ln R)³ [SH helper] | ln R (SH) | (ln R)³ (SH) | In fit range | In op range | fit-masked 1/T | fit-masked ln R | T_narrow (°C) | Resid. in op (°C)
1 | 0.5 | 273.65 | 308000 | 0.0036543 | 12.6379 | 316028.6875 | 2.6067 | 2018.4598 | 12.6379 | 2018.4598 | 0 | 0 | 0 | 0 | 0.90327
2 | 1.7 | 274.85 | 293000 | 0.0036383 | 12.5879 | 297091.6081 | 1.3965 | 1994.6318 | 12.5879 | 1994.6318 | 0 | 0 | 0 | 0 | 1.8893
3 | 4.6 | 277.75 | 250000 | 0.0036004 | 12.4292 | 256444.7197 | 2.5779 | 1920.1326 | 12.4292 | 1920.1326 | 0 | 0 | 0 | 0 | 5.0715
4 | 9 | 282.15 | 203000 | 0.0035442 | 12.221 | 206326.7021 | 1.6388 | 1825.2237 | 12.221 | 1825.2237 | 1 | 0 | 0.0035442 | 12.221 | 9.3604
5 | 11.9 | 285.05 | 181000 | 0.0035082 | 12.1063 | 179434.5299 | -0.8649 | 1774.3086 | 12.1063 | 1774.3086 | 1 | 1 | 0.0035082 | 12.1063 | 11.7797 | -0.12027
6 | 14.5 | 287.65 | 160000 | 0.0034764 | 11.9829 | 158697.6022 | -0.814 | 1720.6359 | 11.9829 | 1720.6359 | 1 | 1 | 0.0034764 | 11.9829 | 14.4274 | -0.072594
7 | 18.4 | 291.55 | 134800 | 0.0034299 | 11.8115 | 132541.0872 | -1.6758 | 1647.8603 | 11.8115 | 1647.8603 | 1 | 1 | 0.0034299 | 11.8115 | 18.1896 | -0.21036
8 | 21.6 | 294.75 | 116480 | 0.0033927 | 11.6655 | 114740.3422 | -1.4935 | 1587.4764 | 11.6655 | 1587.4764 | 1 | 1 | 0.0033927 | 11.6655 | 21.4749 | -0.12513
9 | 25.5 | 298.65 | 97540 | 0.0033484 | 11.488 | 96648.531 | -0.91395 | 1516.126 | 11.488 | 1516.126 | 1 | 1 | 0.0033484 | 11.488 | 25.567 | 0.067003
10 | 28.2 | 301.35 | 87400 | 0.0033184 | 11.3783 | 86046.6174 | -1.5485 | 1473.0805 | 11.3783 | 1473.0805 | 1 | 1 | 0.0033184 | 11.3783 | 28.1556 | -0.044382
11 | 31.2 | 304.35 | 77200 | 0.0032857 | 11.2542 | 75808.0962 | -1.803 | 1425.4062 | 11.2542 | 1425.4062 | 1 | 0 | 0.0032857 | 11.2542 | 31.1367
12 | 35.4 | 308.55 | 65080 | 0.003241 | 11.0834 | 63751.0702 | -2.042 | 1361.4942 | 11.0834 | 1361.4942 | 1 | 0 | 0.003241 | 11.0834 | 35.3371
13 | 38.6 | 311.75 | 56560 | 0.0032077 | 10.9431 | 56044.287 | -0.9118 | 1310.4366 | 10.9431 | 1310.4366 | 1 | 0 | 0.0032077 | 10.9431 | 38.8759
14 | 42.2 | 315.35 | 49300 | 0.0031711 | 10.8057 | 48633.7912 | -1.3513 | 1261.7004 | 10.8057 | 1261.7004 | 0 | 0 | 0 | 0 | 42.4202
15 | 45.5 | 318.65 | 43000 | 0.0031382 | 10.669 | 42825.2847 | -0.40631 | 1214.411 | 10.669 | 1214.411 | 0 | 0 | 0 | 0 | 46.0285
16 | 50 | 323.15 | 36120 | 0.0030945 | 10.4946 | 36157.1365 | 0.10281 | 1155.8405 | 10.4946 | 1155.8405 | 0 | 0 | 0 | 0 | 50.7514
17 | 52.6 | 325.75 | 32800 | 0.0030698 | 10.3982 | 32858.6289 | 0.17875 | 1124.2748 | 10.3982 | 1124.2748 | 0 | 0 | 0 | 0 | 53.4237
18 | 60 | 333.15 | 24890 | 0.0030017 | 10.1222 | 25232.2723 | 1.3751 | 1037.1164 | 10.1222 | 1037.1164 | 0 | 0 | 0 | 0 | 61.3217
19 | 68 | 341.15 | 19400 | 0.0029313 | 9.873 | 19211.49 | -0.9717 | 962.3901 | 9.873 | 962.3901 | 0 | 0 | 0 | 0 | 68.7891
20 | 74 | 347.15 | 15000 | 0.0028806 | 9.6158 | 15788.6713 | 5.2578 | 889.1131 | 9.6158 | 889.1131 | 0 | 0 | 0 | 0 | 76.8551
B-Parameter Model: R(T) = R₀·exp(B·(1/T − 1/T₀)) → plot ln(R) vs 1/T gives slope = B
Slope B (K) | 3872.9814
R₀ fitted at T₀ (Ω) | 98773.4599
Intercept (ln R₀ − B/T₀) | -1.4895
R² of ln(R) vs 1/T | 0.99957
Points used (N) | 20
R @ 25 °C (from fit) (Ω) | 98773.4599
R @ 0 °C (from fit) (Ω) | 324323.0509
R @ 50 °C (from fit) (Ω) | 36157.1365
Steinhart-Hart Model: 1/T = A + B·ln(R) + C·(ln R)³ (LINEST multi-variable regression)
Coefficient A | 0.000826
Coefficient B | 0.00019874
Coefficient C | 1.5683e-07
T @ R=R₀ (SH, °C) check | 25.3369
Narrow-range analysis — fit only the operating-range data (added in v2)
Fit range — T_min (°C): | 5
Fit range — T_max (°C): | 40
Operating range — T_min (°C): | 10
Operating range — T_max (°C): | 30
Narrow-range B-model fit (closed-form least squares on masked rows)
B value (K) | 3816.5643
Intercept (ln R₀ − B/T₀) | -1.2885
R₀ fitted at 25 °C (Ω) | 99939.0133
Points used in fit (N) | 10
R² of narrow fit | 0.99965
Residual statistics — restricted to operating range only (what you'll actually experience)
Points in operating range | 6
Residual RMS (°C) | 0.11979
Residual min (°C) | -0.21036
Residual max (°C) | 0.067003
Residual mean offset (°C) | -0.084289
Use the narrow-range B and intercept above in your firmware. T(R) = B / (ln R − intercept) − 273.15. Residual RMS shows the actual precision you can expect in the operating window.

## Thermistor_2
NTC Thermistor Characterization — NTC-02
Thermistor ID: | NTC-02
Reference T₀ (°C): | 25 | T₀ (K): | 298.15
Measurement Data (fill the blue cells — T_ref is your reference probe, R_meas is multimeter reading)
Point # | T_ref (°C) | T_ref (K) | R_meas (Ω) | 1/T (1/K) | ln(R) | R_fit B-model (Ω) | Residual (%) | (ln R)³ [SH helper] | ln R (SH) | (ln R)³ (SH) | In fit range | In op range | fit-masked 1/T | fit-masked ln R | T_narrow (°C) | Resid. in op (°C)
1 | 0.5 | 273.65 | 312000 | 0.0036543 | 12.6508 | 312531.2853 | 0.17028 | 2024.6488 | 12.6508 | 2024.6488 | 0 | 0 | 0 | 0 | 0.60824
2 | 1.6 | 274.75 | 280000 | 0.0036397 | 12.5425 | 295382.3873 | 5.4937 | 1973.1359 | 12.5425 | 1973.1359 | 0 | 0 | 0 | 0 | 2.7474
3 | 4.9 | 278.05 | 240000 | 0.0035965 | 12.3884 | 250046.8648 | 4.1862 | 1901.2755 | 12.3884 | 1901.2755 | 0 | 0 | 0 | 0 | 5.853
4 | 9.3 | 282.45 | 200000 | 0.0035404 | 12.2061 | 201450.0926 | 0.72505 | 1818.5609 | 12.2061 | 1818.5609 | 1 | 0 | 0.0035404 | 12.2061 | 9.6176
5 | 12 | 285.15 | 179400 | 0.0035069 | 12.0974 | 177014.6773 | -1.3296 | 1770.4075 | 12.0974 | 1770.4075 | 1 | 1 | 0.0035069 | 12.0974 | 11.9108 | -0.089241
6 | 14.6 | 287.75 | 158300 | 0.0034752 | 11.9722 | 156648.7999 | -1.0431 | 1716.0385 | 11.9722 | 1716.0385 | 1 | 1 | 0.0034752 | 11.9722 | 14.597 | -0.0030391
7 | 18.4 | 291.55 | 134000 | 0.0034299 | 11.8056 | 131537.1512 | -1.8379 | 1645.3703 | 11.8056 | 1645.3703 | 1 | 1 | 0.0034299 | 11.8056 | 18.2543 | -0.14574
8 | 21.6 | 294.75 | 116250 | 0.0033927 | 11.6635 | 113937.8801 | -1.9889 | 1586.6696 | 11.6635 | 1586.6696 | 1 | 1 | 0.0033927 | 11.6635 | 21.4469 | -0.15309
9 | 25.2 | 298.35 | 99000 | 0.0033518 | 11.5029 | 97294.8371 | -1.7224 | 1522.016 | 11.5029 | 1522.016 | 1 | 1 | 0.0033518 | 11.5029 | 25.1411 | -0.058897
10 | 27.8 | 300.95 | 88600 | 0.0033228 | 11.3919 | 87012.4422 | -1.7918 | 1478.3832 | 11.3919 | 1478.3832 | 1 | 1 | 0.0033228 | 11.3919 | 27.7483 | -0.051669
11 | 31.6 | 304.75 | 75900 | 0.0032814 | 11.2372 | 74160.6122 | -2.2917 | 1418.963 | 11.2372 | 1418.963 | 1 | 0 | 0.0032814 | 11.2372 | 31.4597
12 | 35.8 | 308.95 | 63550 | 0.0032368 | 11.0596 | 62437.5125 | -1.7506 | 1352.7457 | 11.0596 | 1352.7457 | 1 | 0 | 0.0032368 | 11.0596 | 35.8344
13 | 39.1 | 312.25 | 55200 | 0.0032026 | 10.9187 | 54719.389 | -0.87067 | 1301.7122 | 10.9187 | 1301.7122 | 1 | 0 | 0.0032026 | 10.9187 | 39.3947
14 | 42.6 | 315.75 | 48000 | 0.0031671 | 10.779 | 47716.9633 | -0.58966 | 1252.3627 | 10.779 | 1252.3627 | 0 | 0 | 0 | 0 | 43.0092
15 | 46.2 | 319.35 | 42700 | 0.0031314 | 10.662 | 41578.1512 | -2.6273 | 1212.0218 | 10.662 | 1212.0218 | 0 | 0 | 0 | 0 | 46.1
16 | 50.8 | 323.95 | 34810 | 0.0030869 | 10.4577 | 35025.0472 | 0.61777 | 1143.6774 | 10.4577 | 1143.6774 | 0 | 0 | 0 | 0 | 51.6442
17 | 54 | 327.15 | 31300 | 0.0030567 | 10.3514 | 31174.3566 | -0.40142 | 1109.1593 | 10.3514 | 1109.1593 | 0 | 0 | 0 | 0 | 54.6055
18 | 61.5 | 334.65 | 23509 | 0.0029882 | 10.0651 | 23935.2636 | 1.8132 | 1019.6691 | 10.0651 | 1019.6691 | 0 | 0 | 0 | 0 | 62.8557
19 | 71.5 | 344.65 | 16700 | 0.0029015 | 9.7232 | 17131.4931 | 2.5838 | 919.2271 | 9.7232 | 919.2271 | 0 | 0 | 0 | 0 | 73.2738
20 | 78 | 351.15 | 13500 | 0.0028478 | 9.5104 | 13925.9025 | 3.1548 | 860.2061 | 9.5104 | 860.2061 | 0 | 0 | 0 | 0 | 80.0866
B-Parameter Model: R(T) = R₀·exp(B·(1/T − 1/T₀)) → plot ln(R) vs 1/T gives slope = B
Slope B (K) | 3857.2697
R₀ fitted at T₀ (Ω) | 98142.3055
Intercept (ln R₀ − B/T₀) | -1.4432
R² of ln(R) vs 1/T | 0.9994
Points used (N) | 20
R @ 25 °C (from fit) (Ω) | 98142.3055
R @ 0 °C (from fit) (Ω) | 320700.1506
R @ 50 °C (from fit) (Ω) | 36072.8589
Steinhart-Hart Model: 1/T = A + B·ln(R) + C·(ln R)³ (LINEST multi-variable regression)
Coefficient A | 0.00084548
Coefficient B | 0.00019538
Coefficient C | 1.7029e-07
T @ R=R₀ (SH, °C) check | 25.3729
Narrow-range analysis — fit only the operating-range data (added in v2)
Fit range — T_min (°C): | 5
Fit range — T_max (°C): | 40
Operating range — T_min (°C): | 10
Operating range — T_max (°C): | 30
Narrow-range B-model fit (closed-form least squares on masked rows)
B value (K) | 3820.8243
Intercept (ln R₀ − B/T₀) | -1.3062
R₀ fitted at 25 °C (Ω) | 99601.9652
Points used in fit (N) | 10
R² of narrow fit | 0.9997
Residual statistics — restricted to operating range only (what you'll actually experience)
Points in operating range | 6
Residual RMS (°C) | 0.098985
Residual min (°C) | -0.15309
Residual max (°C) | -0.0030391
Residual mean offset (°C) | -0.083613
Use the narrow-range B and intercept above in your firmware. T(R) = B / (ln R − intercept) − 273.15. Residual RMS shows the actual precision you can expect in the operating window.

## Comparison
Summary & Comparison
Parameter | Thermistor 1 | Thermistor 2 | Ratio / Note
ID | NTC-01 | NTC-02
B value (K) | 3872.9814 | 3857.2697 | 0.99594
R₀ fitted (Ω) at 25 °C | 98773.4599 | 98142.3055 | 0.99361
R² | 0.99957 | 0.9994
R @ 0 °C (Ω) | 324323.0509 | 320700.1506 | 0.98883
R @ 50 °C (Ω) | 36157.1365 | 36072.8589 | 0.99767
SH Coefficient A | 0.000826 | 0.00084548
SH Coefficient B | 0.00019874 | 0.00019538
SH Coefficient C | 1.5683e-07 | 1.7029e-07
Points used (N) | 20 | 20
Narrow-range fit results (v2) — the values you should actually use
Parameter | Thermistor 1 | Thermistor 2 | Ratio / Note
B value (K) — narrow fit | 3816.5643 | 3820.8243 | 1.0011
Intercept — narrow fit | -1.2885 | -1.3062
R₀ at 25 °C (Ω) — narrow | 99939.0133 | 99601.9652 | 0.99663
Points used in fit | 10 | 10
R² of narrow fit | 0.99965 | 0.9997
Points in op range | 6 | 6
Residual RMS in op (°C) | 0.11979 | 0.098985
Residual min in op (°C) | -0.21036 | -0.15309
Residual max in op (°C) | 0.067003 | -0.0030391
Residual mean in op (°C) | -0.084289 | -0.083613